The accuracy of an automated blood pressure measuring device depends on the cuff; the transducer that records the cuff pressure and the amplifier and signal processing (e.g. some form of filter) that records the cuff pressure; the digitisation and subsequent digital signal processing that extract the oscillometric pulses; interpolation between sampled pulses and the algorithm used to analyse the oscillometric pulses and their variation with cuff pressure.  The rate of change of cuff pressure during the measurement phase will affect measurement accuracy (traditionally during linear cuff deflation, but step deflation and measurement during inflation are also used). The inflation of the cuff (assuming measurement during cuff deflation) can also affect accuracy, for example dependent on the difference between the pressure that the cuff is inflated to and the systolic pressure.

Figure:  Simplified block diagram of an automated non-invasive blood pressure measuring device

Please use block capitals to complete the form below and return it together with operating manuals for both the original and derivative device to jackie@bihsoc.org

Manufacturer and Model Details

| Company Name |  |
| --- | --- |
| Derivative Device Model Name/Number |  |
| Original Device Model Name/Number |  |

Validation Study Details of the Original Device

| Protocol used for the study |  |
| --- | --- |
| Author(s) of published study results |  |
| Title of published paper |  |
| Title of Journal/Publication |  |
| Year, Volume, Pages |  |

In order for the BIHS to list a device as equivalent on its website, the manufacturer must state, with reference to the block diagram, if there are any changes in any of the numbered functional blocks

|  | Any change yes / no | Specify |
| --- | --- | --- |
| Cuff characteristics i.e. shape, size or materials |  |  |
| Transducer, amplifier and any signal processing carried out prior to digitisation |  |  |
| Cuff inflation |  |  |
| Cuff deflation |  |  |
| Digitisation (sampling frequency and number of bits), and digital signal processing |  |  |
| Interpolation |  |  |
| Algorithm |  |  |

In addition to the changes notified above, any other changes to the device should be listed below including changes to the casing, memory capacity etc.  Please include photographs of both the original and new device including cuffs. Cuffs should be fully opened out and laid side by side with the original, showing front and back views, preferably with a ruler alongside showing the measurement.

| Change from the original e.g. memory | Details of original device e.g. 30 readings | Change in new device e.g. 50 readings |
| --- | --- | --- |

Signed______________________ (Co. Director)         On behalf of ____________________ (Co.Name)

Date ________________________

Invoicing Details

Please give name and address to whom an invoice should be sent for the BIHS administration fee.

Name:

Address:

# Text boxes

British & Irish Hypertension Society, Blood Pressure Measurement Working Party

Application for Listing a Derivative of a Validated BP Monitor

1

Cuff

Safety cuff deflate system

3

Cuff inflation and rate

Displays and operator controls

- Initiation of measurement

- Display of measured pressures

- Store and recall of measured pressures

- Print out and connectivity

4

Cuff deflation rate

2

Transducer Amplifier and signal processing

5

Digitisation and digital signal processing

6

Interpolation

7

Algorithm

British & Irish Hypertension Society

Blood Pressure Measurement Working Party

Application for Listing a Derivative of a Validated BP Monitor